## Risk Register
Medical Device Cybersecurity Risk Assessment
Cyber issue → C/I/A impact → inherent risk → mitigating control → hazardous situation → patient-safety impact → residual risk → recommendation → standards. Severity & Likelihood 1–5; Risk = Sev×Lik (live formulas). Illustrative — adapt to your device.
ID | Cyber issue | C/I/A impact | Inh. Sev | Inh. Lik | Inherent risk | Inherent level | Mitigating control | Hazardous situation | Patient-safety impact | Res. Sev | Res. Lik | Residual risk | Residual level | Recommendation to remediate | Standards
ra-01 | Unauthorized modification of infusion program / drug library | I, A | 5 | 3 |  |  | Mutual authentication, signed configuration, server-side order validation, dose range limits | Over- or under-infusion of medication | Overdose/underdose — serious harm or death | 5 | 1 |  |  | Enforce signed config + order validation; segment clinical network; alert on anomalous programming | ISO 14971; AAMI SW96; FDA Premarket 2025
ra-02 | Unauthenticated wireless (BLE) command to insulin pump | I | 5 | 3 |  |  | Bonded pairing, authenticated commands, max-dose sanity limits | Incorrect insulin delivery | Hypo-/hyperglycemia — serious harm | 5 | 1 |  |  | Require authenticated pairing; cap maximum dose; alert on anomalous commands | ISO 14971; FDA Premarket 2025
ra-03 | Unauthenticated RF telemetry to implant (pacemaker/ICD) | I, C | 5 | 2 |  |  | Authenticated/encrypted telemetry, staged signed firmware | Altered pacing/defibrillation settings | Inappropriate therapy — death | 5 | 1 |  |  | Deploy authenticated telemetry; monitored firmware update via clinic | ISO 14971; IEC 81001-5-1
ra-04 | Network flood / ransomware / RF jamming (availability loss) | A | 4 | 4 |  |  | Safe-state design, local fallback/buffering, resiliency, offline backups | Therapy interruption / device offline | Delayed therapy or monitoring — harm | 4 | 2 |  |  | Design fail-safe + local autonomy; segment networks; tested backups/restore | ISO 14971; FDA Postmarket; AAMI TIR97
ra-05 | Manipulated or suppressed clinical alert (monitor/CGM) | I, A | 4 | 3 |  |  | Integrity checks, independent on-device alerting, validation | Missed or false clinical alert | Undetected deterioration — harm | 4 | 2 |  |  | Independent local alerting; integrity-verify alert path; anomaly detection | ISO 14971; AAMI TIR57
ra-06 | Adversarial/poisoned input or model-integrity loss (AI) | I | 4 | 3 |  |  | Adversarial testing, input validation, data/model integrity, monitoring | AI misclassification (false negative) | Missed diagnosis — delayed treatment | 4 | 2 |  |  | Adversarial + OOD testing; sign datasets/models; monitor real-world performance | GMLP; NIST AI RMF; ISO 14971
ra-07 | Weak crypto / cloud misconfig / mobile data leakage | C | 2 | 4 |  |  | TLS, encryption at rest, IAM least-privilege, secure mobile storage | Disclosure of PHI | Privacy harm; indirect safety via loss of trust | 2 | 2 |  |  | Enforce TLS + at-rest encryption; CSPM baseline; MASVS mobile storage | ISO 27001; HIPAA; GDPR Art. 32
ra-08 | Unsigned update / no anti-rollback (malicious firmware) | I, A | 5 | 2 |  |  | Signed updates, secure boot, anti-rollback | Malicious firmware installed | Arbitrary device behavior — severe harm | 5 | 1 |  |  | Mandate code signing + secure boot + anti-rollback; verify on receipt | FDA Premarket 2025; IEC 81001-5-1
ra-09 | Unpatched OSS/third-party component (no SBOM/SCA) | C, I, A | 4 | 4 |  |  | SBOM + SCA monitoring, patch SLAs, coordinated disclosure | Exploited known component vulnerability | Compromise — varied patient harm | 4 | 2 |  |  | Maintain SBOM; monitor CISA KEV; meet severity-based patch SLAs | §524B; AAMI TIR97; NIST 800-53
ra-10 | Shared/default/hardcoded credentials | C, I | 4 | 3 |  |  | No default creds, unique authentication, secret scanning | Default/hardcoded credential abuse | Unauthorized control — harm | 4 | 1 |  |  | Remove defaults; provision unique credentials; scan builds for secrets | FDA Premarket 2025; OWASP
ra-11 | Open storage / over-broad IAM in AWS/Azure/GCP | C, A | 3 | 4 |  |  | CIS baseline (CSPM), IAM least-privilege, audit logging | Cloud misconfiguration data exposure | Mass PHI exposure; availability impact | 3 | 2 |  |  | Apply CIS benchmark via CSPM; least-privilege IAM; enable audit logs | ISO 27001; NIST 800-53
ra-12 | Insecure mobile storage / no pinning / rooted device | C, I | 3 | 4 |  |  | Keychain/Keystore, certificate pinning, root/jailbreak detection (MASVS) | Mobile app data leakage / tampering | PHI exposure; manipulated readings | 3 | 2 |  |  | Use platform secure storage + pinning; MASVS testing; integrity checks | OWASP MASVS; FDA Premarket 2025

## Scoring key
Scoring key
Scale | 1 | 2 | 3 | 4 | 5
Severity | Negligible | Minor | Serious | Critical | Catastrophic
Likelihood | Improbable | Remote | Occasional | Probable | Frequent
Risk = Sev×Lik | 1-5 Low | 6-11 Medium | 12-25 High
C/I/A | Confidentiality | Integrity | Availability

## References
References
Standard | Title
ISO 14971:2019 | Application of risk management to medical devices
ISO/TR 24971:2020 | Guidance on the application of ISO 14971
AAMI TIR57:2016 | Principles for medical device security risk management
ANSI/AAMI SW96:2023 | Security risk management for device manufacturers
FDA (2025) | Cybersecurity in Medical Devices — premarket guidance
FDA (2016) | Postmarket Management of Cybersecurity in Medical Devices
IEC 81001-5-1:2021 | Health software security activities in the lifecycle
NIST SP 800-30 Rev.1 | Guide for conducting risk assessments
NIST AI RMF 1.0 | AI risk management (for AI-enabled devices)