## Risk Assessment
Medical Device Risk Assessment (ISO 14971:2019)
Risk ID | Feature / Item / Function | Hazard category | Hazard | Foreseeable sequence of events (cause) | Hazardous situation | Harm | Affected party | S | P1 | P2 | P | Risk Index | Risk Level | Acceptability | Risk control measure(s) | Control type | Control verification (effectiveness) | V&V / Traceability ref. | S' | P1' | P2' | P' | Residual Index | Residual Level | Residual Acceptability | New risk introduced? | Benefit-risk justification | Residual risk disclosed | Status
RA-001 | Infusion pump dosing engine | Software fault | Dose miscalculation after unit change | SW miscomputes rate after unit-of-measure change | Infusion faster than prescribed | Overdose - serious injury | Patient | 4 | 3 | 4 | 3 | 12 | High | Reduce (ALARP) | Dose-rate hard limit + independent watchdog | Protective measure | Fault-injection bench test; watchdog halts within 1 s | SW-VER-218; SRS-4.2 | 4 | 1 | 2 | 1 | 4 | Low | Acceptable | N/A - residual acceptable | No | Yes - IFU Warnings | Closed
RA-002 | Connected device remote interface | Cybersecurity | Unauthorized access (STRIDE: Tampering) | Weak auth lets attacker alter therapy params | Malicious change to therapy settings | Inappropriate therapy - critical injury | Patient | 4 | 3 | 4 | 3 | 12 | High | Reduce (ALARP) | Mutual TLS + signed updates + input validation | Inherently safe design | Pen-test confirms auth-bypass closed | THREAT-MODEL-7; SBOM-3.1 | 4 | 1 | 2 | 1 | 4 | Low | Acceptable | N/A - residual acceptable | No | Yes - Security manual | Closed
RA-003 | Reusable surgical tool | Biological (infection) | Inadequate reprocessing | IFU reprocessing steps unclear to user | Residual bioburden on device | Infection - serious injury | Patient | 3 | 2 | 3 | 2 | 6 | Medium | Acceptable w/ review | Validated reprocessing IFU + user training | Information for safety | Reprocessing validation per ISO 17664 | REP-VAL-09 | 3 | 1 | 2 | 1 | 3 | Low | Acceptable | N/A - residual acceptable | No | Yes - IFU | Closed
RA-004 | Patient monitor alarm subsystem | Functional / loss of performance | Missed critical alarm | Alarm thread starves under high CPU load | Clinician not alerted to deterioration | Delayed treatment - serious injury | Patient | 4 | 2 | 3 | 2 | 8 | Medium | Acceptable w/ review | Watchdog reset + independent audible backup alarm | Protective measure | Load test verifies alarm under max load | SW-VER-104 | 4 | 1 | 2 | 1 | 4 | Low | Acceptable | N/A - residual acceptable | No | Yes - IFU | Closed

## Scales
Scoring Scales & Risk Acceptability
SEVERITY OF HARM (S) |  |  | ACCEPTABILITY (Risk Index = S x P, 1-25) | [5 empty] | Thresholds used by formulas (do not move):
[9 empty] | Intolerable >= | 16
1 | Negligible | No injury or slight inconvenience; no intervention. | 16 | Intolerable | Unacceptable - must reduce risk before proceeding |  |  |  | High >= | 10
2 | Minor | Temporary injury not requiring professional intervention. | 10 | High | Undesirable - reduce as far as possible (ALARP); justify residual |  |  |  | Medium >= | 5
3 | Serious | Injury requiring professional medical intervention. | 5 | Medium | Acceptable with review / documented controls
4 | Critical | Permanent impairment or reversible life-threatening injury. | 1 | Low | Broadly acceptable - no further action required
5 | Catastrophic | Death or permanent loss of life-sustaining function.
PROBABILITY OF HARM (P, via P1 x P2) |  |  | RISK MATRIX (cell = Seve x Prob)
|  |  | Severity\Probability | 1 | 2 | 3 | 4 | 5
1 | Improbable | < 1 in 1,000,000 | 5 | 5 | 10 | 15 | 20 | 25
2 | Remote | 1 in 100k-1M | 4 | 4 | 8 | 12 | 16 | 20
3 | Occasional | 1 in 1k-100k | 3 | 3 | 6 | 9 | 12 | 15
4 | Probable | 1 in 100-1k | 2 | 2 | 4 | 6 | 8 | 10
5 | Frequent | > 1 in 100 | 1 | 1 | 2 | 3 | 4 | 5
P1 = prob. hazardous situation occurs; P2 = prob. it leads to harm. P = ROUNDUP(P1*P2/5,0).

## Hazard Mapping
Hazard Mapping — ISO 14971 / ISO/TR 24971 (Annex C style)
Hazard category | Hazard | Typical hazardous situation | Potential harm | Typ. Severity
Electrical energy | Leakage current / insulation failure | Patient contacts live part | Electric shock / burns | 4
Thermal energy | Surface over-temperature | Skin contact with hot surface | Thermal burn | 3
Mechanical energy | Sharp edge / moving part | Contact during use or service | Laceration / crush injury | 3
Radiation energy | Excess dose / output | Patient over-exposed | Tissue damage / long-term cancer risk | 4
Biological (infection) | Inadequate reprocessing / non-sterile | Contaminated device used on patient | Infection / sepsis | 4
Chemical / biocompatibility | Leachables / cytotoxic material | Prolonged tissue contact | Toxic or allergic reaction | 3
Functional / loss of performance | Therapy delivery failure | Therapy not delivered or incorrect | Injury from missed / incorrect therapy | 4
Software fault | Logic error, crash, race condition | Incorrect output or no output | Misdiagnosis / wrong therapy | 4
Information / labeling | Ambiguous IFU or labeling | User misunderstands intended use | Use error leading to harm | 3
Use error | Confusing UI / alarm fatigue | Critical step omitted or mis-set | Delayed or incorrect treatment | 4
Cybersecurity | Unauthorized access / tampering (STRIDE) | Malicious change to settings or data | Inappropriate therapy / data loss | 4
Environmental / EMC | Electromagnetic interference | Device malfunctions near EM source | Temporary loss of function | 3